## YL09-example
Y520 Conductivity Calibrator
v0.2.1: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | Distilled water | µS/cm |  | 125 | µS/cm | 250 | µS/cm | 500 | µS/cm | 1000 | µS/cm
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values (with default K=1 & B=0)
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date: | 2020-09-10 | [5 empty] | 0.1374 | 23.1475 | 0.25834 | 23.3218 | 0.49886 | 22.734 | 0.98694 | 22.7519
[7 empty] | 0.13618 | 23.1618 | 0.25811 | 23.3283 | 0.49975 | 22.7365 | 0.98677 | 22.7544
1. GET & paste sensor info from Modbus Runner software. | [6 empty] | 0.13693 | 23.1747 | 0.25824 | 23.3322 | 0.49975 | 22.7377 | 0.98677 | 22.757
SN: | YL0918031414 | [5 empty] | 0.13693 | 23.1851 | 0.25824 | 23.3374 | 0.49939 | 22.7403 | 0.98643 | 22.757
Version: | hw 1.2,sw 1.8 | [5 empty] | 0.13621 | 23.1942 | 0.25701 | 23.3413 | 0.49653 | 22.7455 | 0.98643 | 22.7582
Probe: | User Coefficients (previous) | [5 empty] | 0.13621 | 23.2046 | 0.25743 | 23.3439 | 0.49767 | 22.7467 | 0.98612 | 22.7621
K: | 1 | (slope) |  |  |  |  | 0.13667 | 23.2137 | 0.25743 | 23.3465 | 0.49972 | 22.748 | 0.98612 | 22.7634
B: | -0.0021 | (intercept) |  |  |  |  | 0.13665 | 23.2215 | 0.25714 | 23.3479 | 0.4993 | 22.7506 | 0.98632 | 22.7646
2. Re-SET "User Coefficients" to defaults.
K: | 1
B: | 0 |  | Average | #DIV/0! | #DIV/0! |  | 0.13665 | 23.1879 | 0.25774 | 23.3374 | 0.49887 | 22.7424 | 0.98649 | 22.7586
|  |  | Std | #DIV/0! | #DIV/0! |  | 0.00043409 | 0.025661 | 0.0005476 | 0.0093039 | 0.0011738 | 0.0060797 | 0.00030875 | 0.0044831
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! |  | 0.3177% | 0.1107% | 0.2125% | 0.03987% | 0.2353% | 0.02673% | 0.0313% | 0.0197%
Type into pale yellow cells G3, J3, M3, P3 |  |  | Converted units: | #DIV/0! | µS/cm |  | 136.6501 | µS/cm | 257.7447 | µS/cm | 498.8705 | µS/cm | 986.4901 | µS/cm
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | calibrated X | [5 empty] | Measured | Reference
|  |  |  | 136.6501 | 125 |  | 125.6007 | -0.60072 | 125.6007 | [5 empty] | 136.6501 | 125
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 257.7447 | 250 |  | 250.278 | -0.278 | 250.278 | [5 empty] | 257.7447 | 250
Delete green values in pale yellow cells D25:E28 |  |  |  | 498.8705 | 500 |  | 498.5376 | 1.4624 | 498.5376 | [5 empty] | 498.8705 | 500
|  |  |  | 986.4901 | 1000 |  | 1000.5837 | -0.5837 | 1000.5837 | [5 empty] | 986.4901 | 1000
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | 1.0296 | slope |  | slope | 1.0296 | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | -0.015092 | intercept in mS/cm |  | intercept | -15.0923 | µS/cm | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable".
Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
|  |  |  | intercept | -0.015092 | mS/cm
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.

## YL09-blank
Y520 Conductivity Calibrator
v0.2.1: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards |  | µS/cm |  |  | µS/cm |  | µS/cm |  | µS/cm |  | µS/cm
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values (with default K=1 & B=0)
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date:
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
Probe: | User Coefficients (previous)
K: |  | (slope)
B: |  | (intercept)
2. Re-SET "User Coefficients" to defaults.
K: | 1
B: | 0 |  | Average | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  |  | Std | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells G3, J3, M3, P3 |  |  | Converted units: | #DIV/0! | µS/cm |  | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | calibrated X | [5 empty] | Measured | Reference
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | #DIV/0! | [5 empty] | #DIV/0! | 0
5. Delete calibration X:Y pairs that are not valid. |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | #DIV/0! | [5 empty] | #DIV/0! | 0
Delete green values in pale yellow cells D25:E28 |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | #DIV/0! | [5 empty] | #DIV/0! | 0
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | #DIV/0! | [5 empty] | #DIV/0! | 0
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | #DIV/0! | slope |  | slope | #DIV/0! | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | #DIV/0! | intercept in mS/cm |  | intercept | #DIV/0! | µS/cm | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable". Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
|  |  |  | intercept | #DIV/0! | mS/cm
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.