## YL09-example
Y532 pH Calibrator
v0.2.0: 2020-09-10 | [5 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  | Reference Standards | tap water | pH | @ 25°C | 4 | pH | @ 25°C | 6.86 | pH | @ 25°C | 9.18 | pH | @ 25°C
|  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  | pH | mV | °C | pH | mV | °C | pH | mV | °C | pH | mV | °C
Calibration Date: | 2020-09-10 |  |  |  |  | 3.851 | 280.7963 | 20.5838 | 6.5008 | 124.1004 | 20.6961 | 8.8165 | -11.5469 | 20.9232
[6 empty] | 3.8502 | 280.8424 | 20.5835 | 6.542 | 121.6635 | 20.6973 | 8.8162 | -11.5315 | 20.9241
1. GET & paste sensor info from Modbus Runner software. | [5 empty] | 3.8511 | 280.7886 | 20.5844 | 6.5652 | 120.2951 | 20.6979 | 8.8142 | -11.4162 | 20.9235
SN: | YL4318052603 |  |  |  |  | 3.8506 | 280.8194 | 20.5853 | 6.5808 | 119.3726 | 20.7 | 8.8157 | -11.5007 | 20.9235
Version: | hw 1.2,sw 2.0 |  |  |  |  | 3.8502 | 280.8424 | 20.5856 | 6.5905 | 118.796 | 20.7003 | 8.8163 | -11.5392 | 20.9241
pH+NH4_N: | pH coefficients (previous) |  |  |  |  | 3.8483 | 280.9577 | 20.5868 | 6.5997 | 118.2502 | 20.7018 | 8.8182 | -11.6468 | 20.9232
K1: | 102.9719 |  |  |  |  | 3.8421 | 281.3267 | 20.5874 | 6.6065 | 117.8504 | 20.7021 | 8.8188 | -11.6852 | 20.9241
K2: | 94.5699 |  |  |  |  | 3.8439 | 281.2191 | 20.5871 | 6.6117 | 117.543 | 20.7027 | 8.8199 | -11.7467 | 20.9238
K3: | 0.85452 |  |  |  |  | 3.8452 | 281.1422 | 20.5868 | 6.617 | 117.2278 | 20.7036 | 8.8209 | -11.8082 | 20.9244
K4: | 102.9719 |  |  |  |  | 3.8425 | 281.3036 | 20.5871 | 6.6199 | 117.0586 | 20.7048 | 8.8207 | -11.7929 | 20.9244
K5: | 94.6722
K6: | 0.12403 | Average | #DIV/0! | #DIV/0! | #DIV/0! | 3.8475 | 281.0038 | 20.5858 | 6.5834 | 119.2158 | 20.7007 | 8.8178 | -11.6214 | 20.9238
|  | Std | #DIV/0! | #DIV/0! | #DIV/0! | 0.0037138 | 0.22042 | 0.0014717 | 0.037975 | 2.2464 | 0.0028619 | 0.0022915 | 0.13409 | 0.0004568
3. Fill in reference standard known values. |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | 0.09652% | 0.07844% | 0.007149% | 0.5768% | 1.884% | 0.01383% | 0.02599% | -1.154% | 0.002183%
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  | X | Y | [10 empty] | X | Y
Paste into pale yellow cells. |  |  | pH | E(T), mV |  | predicted Y | residual Y | [7 empty] | pH | E(T), mV | temp
|  |  | 4 | 281.0038 | 20.5858 | 280.9376 | 0.066279 | [7 empty] | 4 | 281.0038 | 20.5858
5. Delete calibration X:Y pairs that are not valid. |  |  | 6.86 | 119.2158 | 20.7007 | 119.3637 | -0.14798 | [7 empty] | 6.86 | 119.2158 | 20.7007
Delete green values in pale yellow cells D25:E27 |  |  | 9.18 | -11.6214 | 20.9238 | -11.7031 | 0.081706 | [7 empty] | 9.18 | -11.6214 | 20.9238
6. Re-SET "User Coefficients" to new calibration values. |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | Explanation of pH calibration equations
K1: | 119.2158 |  |  | Slope | mV@pHiso | temp correction | [8 empty] | Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  | Where S0 = 0,E0=0 by default.
K2: | 111.2961 |  | measured 4 to 6.86 | -56.5693 | 111.2961 | -1.715 | [8 empty] | K1=K4 = E (a sensor reading of mV at ph = 6.86)
K3: | -1.715 |  | measured 6.86 to 9.18 | -56.3953 | 111.3204 | -1.9559 | [8 empty] | K2 is actual E0 in acidic solution ph =4 |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K4: | 119.2158 | [13 empty] | K3 is actual S0 in acidic solution ph =4 |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
K5: | 111.3204 |  | theoretical slope = -0.1984*T | -54.193 | [10 empty] | If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
K6: | -1.9559
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.

## YL09-blank
Y511 Conductivity Calibrator
v0.2.0: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards |  | µS/cm |  |  | µS/cm |  | µS/cm |  | µS/cm |  | µS/cm
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date:
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
Probe: | User Coefficients (previous)
K: |  | (slope)
B: |  | (intercept)
2. Re-SET "User Coefficients" to defaults.
K: | 1
B: | 0 |  | Average | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  |  | Std | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells G3, J3, M3, P3 |  |  | Measured | #DIV/0! | µS/cm |  | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | [6 empty] | Measured | Reference
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
5. Delete calibration X:Y pairs that are not valid. |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
Delete green values in pale yellow cells D25:E28 |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | #DIV/0! | slope |  | slope | #DIV/0! | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | #DIV/0! | intercept in mS/cm |  | intercept | #DIV/0! | µS/cm | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable". Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
|  |  |  | intercept | #DIV/0! | mS/cm
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.
[7 empty] | 947.518 | µS/cm | 512.3628 | µS/cm | 218.414 | µS/cm | -0.3745 | µS/cm
[7 empty] | 0 |  | 0 |  | 0
[7 empty] | 0.94752 | 22.3156 | 0.51236 | 22.345 | 0.21841 | 22.3872 | -0.0003745 | 22.3357
[7 empty] | 0.00034983 | 0.031824 | 0.00040162 | 0.023824 | 0.00064849 | 0.030062 | 8.0289e-05 | 0.030177
[7 empty] | 0.03692% | 0.1426% | 0.07839% | 0.1066% | 0.2969% | 0.1343% | -21.44% | 0.1351%
[7 empty] | 0.94757 | 22.3264 | 0.51255 | 22.3103 | 0.21909 | 22.4084 | -0.000359 | 22.3512
[7 empty] | 0.94743 | 22.2681 | 0.51281 | 22.3499 | 0.21792 | 22.3499 | -0.000417 | 22.3499
[7 empty] | 0.94796 | 22.3326 | 0.51189 | 22.3562 | 0.2178 | 22.4146 | -0.000453 | 22.3512
[7 empty] | 0.94712 | 22.3351 | 0.5122 | 22.3636 | 0.21885 | 22.376 | -0.000269 | 22.2904