## YL43-CWS
Y532 pH Calibrator
v0.2.0: 2020-09-10 | [5 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  | Reference Standards | 7 | pH | @ 25°C | 4 | pH | @ 25°C | 6.86 | pH | @ 25°C | 10 | pH | @ 25°C
|  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  | pH | mV | °C | pH | mV | °C | pH | mV | °C | pH | mV | °C
Calibration Date: | 2020-09-08 |  |  | 23.1779 | 22 |  | 191.213 | 22 |  | 26.1069 | 22 |  | -153.528 | 22
193.8498
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
pH+NH4_N: | pH coefficients (previous)
K1: | 26.8142
K2: | 36.8136
K3: | -130.5849
K4: | 26.8142
K5: | 19.5494
K6: | -7.2689 | Average | #DIV/0! | 23.1779 | 22 | #DIV/0! | 192.5314 | 22 | #DIV/0! | 26.1069 | 22 | #DIV/0! | -153.528 | 22
|  | Std | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1.8645 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0.9684% | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  | X | Y | [10 empty] | X | Y
Paste into pale yellow cells. |  |  | pH | E(T), mV |  | predicted Y | residual Y | [7 empty] | pH | E(T), mV | temp
|  |  | 4 | 192.5314 | 22 | 192.0191 | 0.51232 | [7 empty] | 4 | 192.5314 | 22
5. Delete calibration X:Y pairs that are not valid. |  |  | 6.86 | 26.1069 | 22 | 27.0859 | -0.97896 | [7 empty] | 6.86 | 26.1069 | 22
Delete green values in pale yellow cells D25:E27 |  |  | 10 | -153.528 | 22 | -153.9946 | 0.46664 | [7 empty] | 10 | -153.528 | 22
6. Re-SET "User Coefficients" to new calibration values. |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | Explanation of pH calibration equations
K1: | 26.1069 |  |  | Slope | mV@pHiso | temp correction | [8 empty] | Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  | Where S0 = 0,E0=0 by default.
K2: | 17.9603 |  | measured 4 to 6.86 | -58.1904 | 17.9603 | -0.37442 | [8 empty] | K1=K4 = E (a sensor reading of mV at ph = 6.86)
K3: | -0.37442 |  | measured 6.86 to 10 | -57.2086 | 18.0977 | -1.3562 | [8 empty] | K2 is actual E0 in acidic solution ph =4 |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K4: | 26.1069 | [13 empty] | K3 is actual S0 in acidic solution ph =4 |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
K5: | 18.0977 |  | theoretical slope = -0.1984*T | -54.193 | [10 empty] | If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
K6: | -1.3562
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.

## YL43-YosemiTech1
Y532 pH Calibrator
v0.2.0: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | 7 | pH | @ 25°C | 4 | pH | @ 25°C | 6.86 | pH | @ 25°C | 9.18 | pH | @ 25°C
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | pH | mV | °C | pH | mV | °C | pH | mV | °C | pH | mV | °C
Calibration Date: | 2019-06-01 |  |  |  | 23.1779 | 22 |  | 240 | 25 |  | 70 | 25 |  | -60 | 25
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
pH+NH4_N: | pH coefficients (previous)
K1: | 70
K2: | 61.6783
K3: | 0.28056
K4: | 70
K5: | 62.1552
K6: | -3.1255 |  | Average | #DIV/0! | 23.1779 | 22 | #DIV/0! | 240 | 25 | #DIV/0! | 70 | 25 | #DIV/0! | -60 | 25
|  |  | Std | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [10 empty] | X | Y
Paste into pale yellow cells. |  |  |  | pH | E(T), mV |  | predicted Y | residual Y | [7 empty] | pH | E(T), mV | temp
|  |  |  | 4 | 240 | 25 | 238.702 | 1.298 | [7 empty] | 4 | 240 | 25
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 6.86 | 70 | 25 | 72.8981 | -2.8981 | [7 empty] | 6.86 | 70 | 25
Delete green values in pale yellow cells D25:E27 |  |  |  | 9.18 | -60 | 25 | -61.6001 | 1.6001 | [7 empty] | 9.18 | -60 | 25
6. Re-SET "User Coefficients" to new calibration values. |  |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET". | [15 empty] | Explanation of pH calibration equations
K1: | 70 | 0 |  |  | Slope | mV@pHiso | temp correction | [8 empty] | Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  | Where S0 = 0,E0=0 by default.
K2: | 61.6783 | 0 |  | measured 4 to 6.86 | -59.4406 | 61.6783 | 0.28056 | [8 empty] | K1=K4 = E (a sensor reading of mV at ph = 6.86)
K3: | 0.28056 | 7.1054e-15 |  | measured 6.86 to 9.18 | -56.0345 | 62.1552 | -3.1255 | [8 empty] | K2 is actual E0 in acidic solution ph =4 |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K4: | 70 | 0 | [13 empty] | K3 is actual S0 in acidic solution ph =4 |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
K5: | 62.1552 | 0 |  | theoretical slope = -0.1984*T | -54.193 | [10 empty] | If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
K6: | -3.1255 | 0
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.

## YL43-YosemiTech2
Y532 pH Calibrator
v0.2.0: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | 7 | pH | @ 25°C | 4 | pH | @ 25°C | 6.86 | pH | @ 25°C | 9.18 | pH | @ 25°C
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | pH | mV | °C | pH | mV | °C | pH | mV | °C | pH | mV | °C
Calibration Date: | 2020-09-09 |  |  |  | 23.1779 | 22 |  | 232.48 | 25.6 |  | 73.46 | 26 |  | -42.74 | 25.4
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
pH+NH4_N: | pH coefficients (previous)
K1: | 73.46
K2: | 65.6758
K3: | -3.757
K4: | 73.46
K5: | 66.4479
K6: | -9.1532 |  | Average | #DIV/0! | 23.1779 | 22 | #DIV/0! | 232.48 | 25.6 | #DIV/0! | 73.46 | 26 | #DIV/0! | -42.74 | 25.4
|  |  | Std | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [10 empty] | X | Y
Paste into pale yellow cells. |  |  |  | pH | E(T), mV |  | predicted Y | residual Y | [7 empty] | pH | E(T), mV | temp
|  |  |  | 4 | 232.48 | 25.6 | 230.3782 | 2.1018 | [7 empty] | 4 | 232.48 | 25.6
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 6.86 | 73.46 | 26 | 78.1527 | -4.6927 | [7 empty] | 6.86 | 73.46 | 26
Delete green values in pale yellow cells D25:E27 |  |  |  | 9.18 | -42.74 | 25.4 | -45.331 | 2.591 | [7 empty] | 9.18 | -42.74 | 25.4
6. Re-SET "User Coefficients" to new calibration values. |  |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET". | [15 empty] | Explanation of pH calibration equations
K1: | 73.46 | 0 |  |  | Slope | mV@pHiso | temp correction | [8 empty] | Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  | Where S0 = 0,E0=0 by default.
K2: | 65.6758 | 0 |  | measured 4 to 6.86 | -55.6014 | 65.6758 | -3.757 | [8 empty] | K1=K4 = E (a sensor reading of mV at ph = 6.86)
K3: | -3.757 | 0 |  | measured 6.86 to 9.18 | -50.0862 | 66.4479 | -9.1532 | [8 empty] | K2 is actual E0 in acidic solution ph =4 |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K4: | 73.46 | 0 | [13 empty] | K3 is actual S0 in acidic solution ph =4 |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
K5: | 66.4479 | 0 |  | theoretical slope = -0.1984*T | -54.193 | [10 empty] | If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
K6: | -9.1532 | 0
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.