## YL09-example
Y511 Conductivity Calibrator
v0.2.0: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | Distilled water | µS/cm |  | 125 | µS/cm | 250 | µS/cm | 500 | µS/cm | 1000 | µS/cm
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date: | 2018-09-25 |  |  | -0.014417 | 22.3115 |  |  |  | 0.21482 | 22.3587 | 0.44703 | 22.3599 | 0.90414 | 22.3314
|  |  |  | -0.015023 | 22.252 |  |  |  | 0.21499 | 22.4208 | 0.445 | 22.3823 | 0.90412 | 22.3401
1. GET & paste sensor info from Modbus Runner software. |  |  |  | -0.01496 | 22.3152 |  |  |  | 0.21466 | 22.4196 | 0.44487 | 22.3848 | 0.90382 | 22.283
SN: | YL0916112101 |  |  | -0.01496 | 22.3164 |  |  |  | 0.21507 | 22.4208 | 0.44621 | 22.3673 | 0.90432 | 22.3512
Version: | hw 1.2,sw 1.6 |  |  | -0.015127 | 22.3189 |  |  |  | 0.21507 | 22.422 | 0.44517 | 22.391 | 0.90306 | 22.3549
Probe: | User Coefficients (previous) |  |  | -0.015331 | 22.3189 |  |  |  | 0.21477 | 22.4208 | 0.44566 | 22.3711 | 0.90358 | 22.3251
K: | 1 | (slope) |  | -0.015018 | 22.3226 |  |  |  | 0.21337 | 22.422 | 0.44518 | 22.3922 | 0.90297 | 22.3636
B: | 0.0194 | (intercept) |  | -0.015018 | 22.3152 |  |  |  | 0.21322 | 22.4208 | 0.44518 | 22.3959 | 0.90344 | 22.3636
|  |  |  | -0.01546 | 22.3152 |  |  |  | 0.21315 | 22.4233 | 0.44519 | 22.3748 | 0.90418 | 22.3053
2. Re-SET "User Coefficients" to defaults. |  |  |  | -0.015128 | 22.3251 |  |  |  | 0.21379 | 22.3611 | 0.4451 | 22.3972 | 0.90387 | 22.3686
K: | 1
B: | 0 |  | Average | -0.015044 | 22.3111 |  | #DIV/0! | #DIV/0! | 0.21429 | 22.409 | 0.44546 | 22.3816 | 0.90375 | 22.3387
|  |  | Std | 0.00027455 | 0.021137 |  | #DIV/0! | #DIV/0! | 0.00081032 | 0.025904 | 0.00067238 | 0.01286 | 0.00047364 | 0.027909
3. Fill in reference standard known values. |  |  | %CV | -1.825% | 0.09474% |  | #DIV/0! | #DIV/0! | 0.3781% | 0.1156% | 0.1509% | 0.05746% | 0.05241% | 0.1249%
Type into pale yellow cells G3, J3, M3, P3 |  |  | Measured | -15.0442 | µS/cm |  | #DIV/0! | µS/cm | 214.2935 | µS/cm | 445.4595 | µS/cm | 903.7492 | µS/cm
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | [6 empty] | Measured | Reference
[15 empty] | #DIV/0! | 125
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 214.2935 | 250 |  | 249.3731 | 0.62686 | [6 empty] | 214.2935 | 250
Delete green values in pale yellow cells D25:E28 |  |  |  | 445.4595 | 500 |  | 500.9431 | -0.94306 | [6 empty] | 445.4595 | 500
|  |  |  | 903.7492 | 1000 |  | 999.6838 | 0.3162 | [6 empty] | 903.7492 | 1000
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | 1.0883 | slope |  | slope | 1.0883 | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | 0.016165 | intercept in mS/cm |  | intercept | 16.165 | µS/cm | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable". Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
|  |  |  | intercept | 0.016165 | mS/cm
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.
[7 empty] | 947.518 | µS/cm | 512.3628 | µS/cm | 218.414 | µS/cm | -0.3745 | µS/cm
[7 empty] | 0 |  | 0 |  | 0
[7 empty] | 0.94752 | 22.3156 | 0.51236 | 22.345 | 0.21841 | 22.3872 | -0.0003745 | 22.3357
[7 empty] | 0.00034983 | 0.031824 | 0.00040162 | 0.023824 | 0.00064849 | 0.030062 | 8.0289e-05 | 0.030177
[7 empty] | 0.03692% | 0.1426% | 0.07839% | 0.1066% | 0.2969% | 0.1343% | -21.44% | 0.1351%
[7 empty] | 0.94757 | 22.3264 | 0.51255 | 22.3103 | 0.21909 | 22.4084 | -0.000359 | 22.3512
[7 empty] | 0.94743 | 22.2681 | 0.51281 | 22.3499 | 0.21792 | 22.3499 | -0.000417 | 22.3499
[7 empty] | 0.94796 | 22.3326 | 0.51189 | 22.3562 | 0.2178 | 22.4146 | -0.000453 | 22.3512
[7 empty] | 0.94712 | 22.3351 | 0.5122 | 22.3636 | 0.21885 | 22.376 | -0.000269 | 22.2904

## YL09-blank
Y511 Conductivity Calibrator
v0.2.0: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards |  | µS/cm |  |  | µS/cm |  | µS/cm |  | µS/cm |  | µS/cm
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date:
1. GET & paste sensor info from Modbus Runner software.
SN:
Version:
Probe: | User Coefficients (previous)
K: |  | (slope)
B: |  | (intercept)
2. Re-SET "User Coefficients" to defaults.
K: | 1
B: | 0 |  | Average | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  |  | Std | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Type into pale yellow cells G3, J3, M3, P3 |  |  | Measured | #DIV/0! | µS/cm |  | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm | #DIV/0! | µS/cm
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | [6 empty] | Measured | Reference
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
5. Delete calibration X:Y pairs that are not valid. |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
Delete green values in pale yellow cells D25:E28 |  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
|  |  |  | #DIV/0! | 0 |  |  | #VALUE! | [6 empty] | #DIV/0! | 0
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | #DIV/0! | slope |  | slope | #DIV/0! | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | #DIV/0! | intercept in mS/cm |  | intercept | #DIV/0! | µS/cm | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable". Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
|  |  |  | intercept | #DIV/0! | mS/cm
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.
[7 empty] | 947.518 | µS/cm | 512.3628 | µS/cm | 218.414 | µS/cm | -0.3745 | µS/cm
[7 empty] | 0 |  | 0 |  | 0
[7 empty] | 0.94752 | 22.3156 | 0.51236 | 22.345 | 0.21841 | 22.3872 | -0.0003745 | 22.3357
[7 empty] | 0.00034983 | 0.031824 | 0.00040162 | 0.023824 | 0.00064849 | 0.030062 | 8.0289e-05 | 0.030177
[7 empty] | 0.03692% | 0.1426% | 0.07839% | 0.1066% | 0.2969% | 0.1343% | -21.44% | 0.1351%
[7 empty] | 0.94757 | 22.3264 | 0.51255 | 22.3103 | 0.21909 | 22.4084 | -0.000359 | 22.3512
[7 empty] | 0.94743 | 22.2681 | 0.51281 | 22.3499 | 0.21792 | 22.3499 | -0.000417 | 22.3499
[7 empty] | 0.94796 | 22.3326 | 0.51189 | 22.3562 | 0.2178 | 22.4146 | -0.000453 | 22.3512
[7 empty] | 0.94712 | 22.3351 | 0.5122 | 22.3636 | 0.21885 | 22.376 | -0.000269 | 22.2904