## YL43-example
Y532 pH Calibrator
v0.2.1: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | tap water | pH | at 25°C | 4 | pH | @ 25°C | 6.86 | pH | at 25°C | 9.18 | pH | at 25°C
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | pH | E (mV) | T (°C) | pH | E (mV) | T (°C) | pH | E (mV) | T (°C) | pH | E (mV) | T (°C)
Calibration Date: | 2021-05-30 | [5 empty] | 4.1984 | 271.3868 | 23.5913 | 7.136 | 103.5669 | 23.9647 | 9.1663 | -12.1926 | 24.047
[7 empty] | 4.2028 | 271.1331 | 23.5888 | 7.1335 | 103.713 | 23.9608 | 9.1677 | -12.2695 | 24.0477
1. GET & paste sensor info from Modbus Runner software. | [6 empty] | 4.2028 | 271.1331 | 23.5852 | 7.1325 | 103.7668 | 23.9535 | 9.176 | -12.7461 | 24.0489
SN: | YL4318052603 | [5 empty] | 4.202 | 271.1792 | 23.5828 | 7.1293 | 103.9513 | 23.9492 | 9.1833 | -13.1612 | 24.0501
Version: | hw 1.2,sw 2.0 | [5 empty] | 4.1991 | 271.3406 | 23.5776 | 7.1293 | 103.9513 | 23.9386 | 9.1885 | -13.4611 | 24.0519
pH+NH4_N: | pH coefficients (previous) | [5 empty] | 4.2007 | 271.2484 | 23.5746 | 7.1258 | 104.1512 | 23.9341 | 9.1858 | -13.1535 | 24.0532
K1: | 119.2158 | mV at pH= 6.86 |  |  |  |  | 4.2009 | 271.233 | 23.5676 | 7.1244 | 104.228 | 23.9265 | 9.1837 | -13.1843 | 24.0543
K2: | 111.2961 | E0, mV at pH= 7, acidic |  |  |  |  | 4.2009 | 271.233 | 23.5658 | 7.1253 | 104.2203 | 23.9225 | 9.1854 | -13.2842 | 24.0571
K3: | -1.7378 | S0, acidic
K4: | 119.2158 | mV at pH= 6.86
K5: | 111.3204 | E0, mV at pH= 7, alkaline
K6: | -1.9559 | S0, alkaline | Average | #DIV/0! | #DIV/0! | #DIV/0! | 4.2009 | 271.2359 | 23.5792 | 7.1295 | 103.9436 | 23.9438 | 9.1796 | -12.9316 | 24.0513
|  |  | Std | #DIV/0! | #DIV/0! | #DIV/0! | 0.0016013 | 0.091227 | 0.0094412 | 0.0042243 | 0.24698 | 0.015697 | 0.008567 | 0.47651 | 0.0034898
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | 0.03812% | 0.03363% | 0.04004% | 0.05925% | 0.2376% | 0.06556% | 0.09333% | -3.685% | 0.01451%
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y |  |  |  | 2-point | [6 empty] | X | Y
Paste into pale yellow cells. |  |  |  | pH | E (mV) | T (°C) | predicted Y | residual Y | predicted X | [6 empty] | pH | E (mV) | T (°C)
|  |  |  | 4 | 271.2359 | 23.5792 | 268.1428 | 3.0931 | 4 | [6 empty] | 4 | 271.2359 | 23.5792
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 6.86 | 103.9436 | 23.9438 | 110.8498 | -6.9062 | 6.86 | [6 empty] | 6.86 | 103.9436 | 23.9438
Delete green values in pale yellow cells D25:E27 |  |  |  | 9.18 | -12.9316 | 24.0513 | -16.7446 | 3.813 | 9.18 | [6 empty] | 9.18 | -12.9316 | 24.0513
6. Re-SET "User Coefficients" to new calibration values. |  |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET".
K1: | 103.9436 | mV at pH= 6.86 |  |  | S, slope | E0 | S0, slope correction
K2: | 95.7545 | E0, mV at pH= 7, acidic |  |  | mV/pH | mV at pHiso
K3: | -0.45663 | S0, acidic |  | acidic pH 4 to 6.86 | -58.4938 | 95.7545 | -0.45663
K4: | 103.9436 | mV at pH= 6.86 |  | alkaline pH 6.86 to 9.18 | -50.3772 | 96.8908 | -8.5945
K5: | 96.8908 | E0, mV at pH= 7, alkaline
K6: | -8.5945 | S0, alkaline |  | ideal slope: | -58.9504 | mV/pH
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.
Background Information
Explanation of pH calibration equations, from YosemiTech
Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  |  | Where S0 = 0,E0=0 by default.
K1=K4 = E (a sensor reading of mV at ph = 6.86)
K2 is actual E0 in acidic solution ph =4 |  |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K3 is actual S0 in acidic solution ph =4 |  |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
Bier2018-Hach-ElectrochemTheoryPractice_DOC182.53.90629.pdf
https://www.google.com/url?sa=t&rct=j&q=&esrc=s&source=web&cd=&ved=2ahUKEwiEhbGrxN_rAhUMbc0KHXdmAJ8QFjAJegQIBBAB&url=https%3A%2F%2Fwww.hach.com%2Fasset-get.download.jsa%3Fid%3D7672493965&usg=AOvVaw0aMsv1tmXVclWovq6qngYQ
Nerstt equation for pH:
| slope = 2.303*RT/nF = -0.1984*T = | -59.153 | mV/pH at 25°C | -54.193 | Adjustment for °C vs °K
| R = | 8.3145 | J K−1 mol−1
| F = | 96485.332 | C mol−1
| 2.303*R/nF = | 0.19846 | mV/K, where 2.303 converts from ln to log10; 1000 mV/V; 1 C / 1 V
Theory and Practice of pH Measurement, Chapter 3 & Appendix B
https://www.emerson.com/documents/automation/manual-theory-practice-of-ph-measurement-rosemount-en-70736.pdf
https://en.wikipedia.org/wiki/Nernst_equation